Table 1 The Kuduwave (KW) evaluated against the required audiometer specifications

| SANS 8253-1 requirements | Kuduwave complies |  | Evidence |
| --- | --- | --- | --- |
|  | Yes | No |  |
| - Can the KW perform an audiometric test manually? | √ |  | Specifications and observed and references 10 and 11 |
| - Can KW perform an audiometric test with automatic-recording? | √ |  | Specifications and observed |
| - When testing automatically is the order of presentation of test tones from 1,000 Hz upwards, followed by the lower frequency range, in descending order? | √ |  | Specifications and observed |
| - Is a repeat test carried out at 1,000 Hz on the ear tested first? | √ |  | Specifications and observed |
| - Does the KW use either the bracketing or the ascending method of test tones presentation? | √ |  | Specifications and observed both. The shortened ascending method (Hughson-Westlake) is also used4 |
| - Is the KW constructed in accordance with IEC60645-1? | √ |  | Certificates from Intertek |
| - Is the KW at least a Type 4 audiometer? | √ |  | Specifications and observed Type 2 (Type 2 performs better than Type 4) |
| - Can the KW present at least the frequencies 0.5, 1, 2, 3, 4, 6 and 8 kHz tones? | √ |  | Specifications and observed |
| - Can the KW present the 8,000 Hz tone at least at 70 dB? | √ |  | Specifications and observed |
| - Can the KW present all other tones from 0 to 70 dB? | √ |  | Specifications and observed |
| - Can the KW present tones in 5 dB steps? | √ |  | Specifications and observed |
| - Does the KW have dual fixed headphones? | √ |  | Specifications and observed |
| - Does the KW have a stimulus button? | √ |  | Specifications and observed |
| - Does the KW have a patient response button? | √ |  | Specifications and observed |
| - Can the KW present pulsed tone and continuous tones? | √ |  | Specifications and observed |

Table 2 Reference values for frequencies and sound pressure levels, for insert earphones as specified in ISO 389-2, in an acoustic coupler that complies with IEC 60318-5 (Source SANS 10154-1 Table 2)

| Frequency Setting (Hz) | Permissible range of frequencies (Hz) | Nominal sound pressure level at a setting of 70 dB (dB) | Permissible range of sound pressure levels at a setting of 70 dB (dB) |
| --- | --- | --- | --- |
| 250 | 242 – 258 | 84,0 | 81,0 – 87,0 |
| 500 | 485 – 515 | 75,5 | 72,5 – 78,5 |
| 1 000 | 970 – 1 030 | 70,0 | 67,0 – 73,0 |
| 2 000 | 1 940 – 2 060 | 73,0 | 70,0 – 76,0 |
| 3 000 | 2 910 – 3 090 | 73,5 | 70,5 – 76,5 |
| 4 000 | 3 880 – 4 120 | 75,5 | 72,5 – 78,5 |
| 6 000 | 5 820 – 6 180 | 72,0 | 67,0 – 77,0 |
| 8 000 | 7 760 – 8 240 | 70,0 | 65,0 – 75,0 |

Table 4 The Kuduwave (KW) evaluated against the calibration requirements

| SANS 10154-1:2012 | Kuduwave complies |  | Evidence |
| --- | --- | --- | --- |
|  | Yes | No |  |
| - Does the KW calibration process comply with calibration requirements for ISO389-2 and IEC60318-5? | √ |  | Accreditation letter from UK-based accreditation company Intertek (affiliated to IEC). |
| - Is KW calibration done at an organisation that has personnel with the necessary education, training, technical knowledge and experience for their assigned functions (at least one Registered Professional Engineer, Scientist or Technologist)? | √ |  | Kuduwave manufacturers calibrate the equipment. Staff comprises technologists. |
| - Is KW calibration done at an organisation that has the necessary facilities with properly maintained and calibrated instruments? | √ |  | Accredited by LRQ for management system in compliance with ISO 13485:2003. |
| - Is KW calibration done at an organisation that implements and maintains a quality management system? | √ |  | Observation of organisational facilities. ISO 13485:2003 documentation. |
| - Does KW recommend that if the instrument is used for mobile purposes it should be calibrated where the maximum interval between such checks should not exceed 3 months? | √ |  | Recommended on invoices and calibration certificates. |
| - Does the KW calibration process use an occluded ear simulator or an acoustic coupler that complies with ISO 389-2 & IEC 70318-5? | √ |  | Laboratory standard operating procedure documents. |
| - Do KW calibration checks include measuring the sound pressure levels from the insert earphones in an acoustic coupler or ear simulator and comparing results with SANS 8253-1? | √ |  | Laboratory standard operating procedure documents. |
| - Is SANS 10154-1 Table 2, entitled Reference values for frequencies and sound pressure levels, for insert earphones as specified in ISO 389-2, in an acoustic coupler that complies with IEC 60318-5, used as the reference values for frequencies and sound pressure levels calibrating the KW earphones? | √ |  | Laboratory standard operating procedure documents. |
| - Do owners get a printout that compares their machine with Table 2 and 3 of SANS 10154-1? | √ |  | Calibration certificate states that this information is available on request |
| - Do KW calibration checks include measuring of the frequencies of test signals and comparing results with SANS 8253-1? | √ |  | Laboratory standard operating procedure documents. |
| - Does the recommended maximum interval between calibration checks not exceed 12 months? | √ |  | Calibration certificates. |
| - During calibration of the KW is a class 1 sound level meter used that has a pressure-calibrated condenser microphone suitable for the ear simulator and that complies with the IEC 61672-1 standard? | √ |  | Laboratory standard operating procedure documents. |
| - During calibration of the KW is a sound level meter used that has a one-third-octave-band filter set that complies with the IEC 61260 standard? | √ |  | Laboratory standard operating procedure documents. |
| - During calibration of the KW are ear simulators or acoustic couplers used that comply with IEC 60318-1[4], IEC 60318-3[5], IEC 60318-4[6], and IEC 60318-5[7]? | √ |  | Laboratory standard operating procedure documents. |
| - During calibration of the KW is a digital frequency counter used? | √ |  | Laboratory standard operating procedure documents. |
| - During calibration of the KW is an oscilloscope used? | √ |  | Laboratory standard operating procedure documents. |
| - During calibration of the KW is a contact thermometer for checking the operating temperature (23°C) of the mechanical coupler used? | √ |  | Laboratory standard operating procedure documents. |
| - After calibration of the KW is a calibration label attached to the equipment, which gives the date on which the next objective test is due? | √ |  | The Kuduwave software prompts the user. Calibration certificate example includes date. |
| - If after completion of the calibration procedure the audiometer performance is found to be in accordance with the relevant standards, is a certificate issued that is valid for a period of one year from the date of calibration? | √ |  | Calibration certificate example. |
| - Does the certificate state that it will immediately become invalid if either the audiometer or its earphones or inserts are subjected to any misuse or rough handling, or to repairs or replacement? | √ |  | Calibration certificate example. |
| - Does the certificate state that it will immediately become invalid if moved from the site of calibration by road, rail or air, unless the audiometer shall be mounted on an anti-vibration platform whenever it is moved? | √ |  | Calibration certificate example. |
| - Does the calibration certificate state that the air conduction calibration of the audiometer has been checked in accordance with SANS 10154 and has been found to be in agreement with the recommended limits? | √ |  | Calibration certificate example. |
| - Does the calibration certificate state the date of calibration? | √ |  | Calibration certificate example. |
| - Does the calibration certificate state the name and address of the calibrating laboratory/organisation? | √ |  | Calibration certificate example. |
| - Does the calibration certificate state the make, model number and serial number of the audiometer? | √ |  | Calibration certificate example. |
| - Does the calibration certificate state the type and serial number of the earphones or inserts? | √ |  | Calibration certificate example. |
| - Does the calibration certificate state the calibration certificate numbers and dates of calibration of all equipment used to check the audiometer? | √ |  | Calibration certificate example. |
| - Does the calibration certificate contain the name and signature of the person conducting the calibration? | √ |  | Calibration certificate example. |

Table 5 Maximum permissible ambient sound pressure levels, Lperm, in octave bands for air conduction audiometry for typical supra-aural headsets (Source: SANS 10182 Table 1)

| Mid-frequency of octave band Hz | Maximum permissible ambient sound pressure levels Lperm dB |  |
| --- | --- | --- |
|  | Diagnostic audiometry | Screening audiometry |
| 125 | 29,0 | 52,0 |
| 250 | 21,0 | 38,5 |
| 500 | 20,5 | 22,0 |
| 1 000 | 24,0 | 24,0 |
| 2 000 | 31,0 | 31,0 |
| 4 000 | 37,0 | 37,0 |
| 8 000 | 35,5 | 35,5 |

Table 6 Calculating the maximum permissible ambient noise levels for use of Kuduwave in a room to test accurately to 0 dBHL

| Hz | Average attenuation provided by TDH39 | Average attenuation provided by Kuduwave (UP study) | Difference between the average attenuation provided by two earphones | Permissible ambient noise levels for screening audiometry | Maximum permissible ambient noise levels when an Ambidome is used |
| --- | --- | --- | --- | --- | --- |
| 125 | 3dB | 38 dB | 34 dB | 52 dB | 52+34=86 dB |
| 250 | 5 dB | 41 dB | 35 dB | 38 dB | 38+35=73 dB |
| 500 | 7 dB | 46 dB | 38 dB | 22 dB | 22+38 =60 dB |
| 1 000 | 15 dB | 47 dB | 32 dB | 24 dB | 24+32=56 dB |
| 2 000 | 26 dB | 50 dB | 24 dB | 31 dB | 31+24=55 dB |
| 4 000 | 31 dB | 60 dB | 28 dB | 37 dB | 37+28=65 dB |
| 8 000 | 24 dB | 53 dB | 28 dB | 35 dB | 35+28=63 dB |